## Worksheet
OC4000
[5 empty] | Save As Before Closing!!!! | [10 empty] | OC3000
| LINEAR AXES TESTING |  |  |  |  | 2
|  | X Axis | [5 empty] | Y Axis
|  | Repeatability: |  | Reticle Scale | X |  |  | Repeatability: |  | Reticle Scale | Y
| Start Pt. |  | Linear Displacement: |  |  |  | Start Pt. |  | Linear Displacement:
| On Std. | Results | ACTUAL | Inc. Dev. | Abs. Dev. |  | On Std. | Results | ACTUAL | Inc. Dev. | Abs. Dev.
1 | 1 |  | -5.4e-05 |  |  | 0 | 1 |  | -2.7e-05 |  |  |  |  | 'OC3000 Values'!$C$6:$F$30 | 'OC3000 Values'!$H$6:$K$18
2 | 2 |  | -7.2e-05 |  |  |  | 2 |  | -4.2e-05 |  |  | Index (H3) | Values | X | Y
3 | 3 |  | -0.000158 |  |  |  | 3 |  | -6.4e-05 |  |  | 1 | OC4000 | 'OC4000 Values'!$C$6:$F$30 | 'OC4000 Values'!$H$6:$K$18
4 | 4 |  | -0.000191 |  |  |  | 4 |  | -5.2e-05 |  |  | 2 | OC3000 | 'OC3000 Values'!$C$6:$F$30 | 'OC3000 Values'!$H$6:$K$18
5 | 5 |  | -0.000194 |  |  |  | 5 |  | -6.4e-05 |  |  | 3 | OC1000 | 'OC4000 Values'!$C$6:$F$30
6 | 6 |  | -0.000231 |  |  |  | 6 |  | -2.7e-05
7 | 7 |  | -0.000196 |  |  |  | 7
8 | 8 |  | -0.000247 |  |  |  | 8
9 | 9 |  | -0.000251 |  |  |  | 9
10 | 10 |  | -0.000327 |  |  |  | 10
11 | 11 |  | -0.000338 |  |  |  | 11
12 | 12 |  | -0.000372 |  |  |  | 12
|  |  | Max Inc. Dev. | 0 |  |  |  |  | Max Inc. Dev. | 0
|  |  | Max Abs. Dev. | 0 |  |  |  |  | Max Abs. Dev. | 0
|  | Max | 0 | 0 |  |  |  | Max | 0 | 0
|  | Min | 0 | 0 |  |  |  | Min | 0 | 0
|  | Abs Min | 0 | 0 |  |  |  | Abs Min | 0 | 0
| MAGINIFICATION TESTING | [6 empty] | Angle Testing
|  | Setting used on Magscale: | 10 | X |  |  |  |  | Standard Angle Values |  | Protractor Results: |  | (Decimal Degrees can be entered in the "DEG" column) |  |  | OC1000 | [5 empty] | Dev
|  | Comparator Magnification: | 10 | X |  |  |  | Nominal | Actual | Dev. | Deg | Min | Sec | Deviation (DMS) | Deviation (decimal round as needed) |  |  | Std Cnts | UUT Cnts | Difference |  | Deg | Min | Sec | dec deg dev
Circle Measured: |  | Magnified Standard Diameter |  | Reading: | Deviation | Tolerance |  | 15 | 14.9989 | -0.0011 |  |  |  |  | 0 |  | 0 | 53996.04 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0
B |  | 2.4989 | [5 empty] | 30 | 29.9967 | -0.0033 |  |  |  |  | 0 |  | 0 | 107988.12 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0
C |  | 6.2488 | [5 empty] | 45 | 44.9963 | -0.0037 |  |  |  |  | 0 |  | 0 | 161986.68 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0
F |  | 0.6239 | [5 empty] | 60 | 59.9962 | -0.0038 |  |  |  |  | 0 |  | 0 | 215986.32 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0
H |  | 2.4991 | [5 empty] | 75 | 74.9969 | -0.0031 |  |  |  |  | 0 |  | 0 | 269988.84 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0
|  | #N/A | [5 empty] | 90 | 89.9996 | -0.0004 |  |  |  |  | 0 |  | 0 | 323998.56 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0
|  | #N/A | [5 empty] | 105 | 104.997 | -0.003 |  |  |  |  | 0 |  | 0 | 377989.2 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0
|  | #N/A | [5 empty] | 120 | 119.9963 | -0.0037 |  |  |  |  | 0 |  | 0 | 431986.68 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0
|  | #N/A | [5 empty] | 135 | 134.9975 | -0.0025 |  |  |  |  | 0 |  | 0 | 485991 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0
|  | #N/A | [5 empty] | 150 | 149.9988 | -0.0012 |  |  |  |  | 0 |  | 0 | 539995.68 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0
|  | #N/A | [5 empty] | 165 | 164.9941 | -0.0059 |  |  |  |  | 0 |  | 0 | 593978.76 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0
|  | #N/A | [5 empty] | 180 | 179.9966 | -0.0034 |  |  |  |  | 0 |  | 0 | 647987.76 | 0 | 0 | -1 | 0 | 0 | 0 | 0 | 0
|  | #N/A | [14 empty] | 0

## Machine info
Report No. | RESULTS
Customer |  | X AXIS
Item |  | L.D.A Incremental
Model No. |  | L.D.A Maximum
Unit No. |  | Repeatability
Serial No.
Location
Inspector |  | Y AXIS
Date Cal. |  | L.D.A Incremental
Date Due |  | L.D.A Maximum
Temp. |  | Repeatability
Humidity
X Axis Range
Y Axis Range |  | Magnification Error(Avg)
Scale Res. |  | Protractor Error(Max)
Protractor Res.
Lens Mag.

## OC3000 Values
[7 empty] | Due Date |  | 2027-06-29
|  |  | GAGE LINE RETICLE CERTIFICATION VALUES |  |  |  | Calibration Date |  | 2024-06-29
OC3000 VALUES
Top Inch Scale |  |  |  | Vertical Inch Scale |  |  |  | CIRCLES | ACTUAL SIZE |  | ANGLES |  |  | Standard Counts |  |  |  | minutes extracted | Extraction |  |  |  |  | seconds extracted
Nominal | error (µ") | Actual |  | Nominal | error (µ") | Actual |  | A | 0.06241 |  | Nominal | ACTUAL VALUES | Dev. | Deg | Min | Sec | Std Counts |  | Deg | Min | seconds extractor |  |  |  | Sec
0 |  | 0 | 0 | 0 |  | 0 | 0 | B | 0.24989 |  | 15 | 14.9989 | -0.0011 | 14 | 59 | 56.04 | 53996.04 | 0.9989 | 14 | 59.934 | 59 | 0.066 | 1 | 0.934 | 56.04
0.5 |  | 0.50001 | 1.4e-05 | 0.5 | 17 | 0.50002 | 1.7e-05 | C | 0.62488 |  | 30 | 29.9967 | -0.0033 | 29 | 59 | 48.12 | 107988.12 | 0.9967 | 29 | 59.802 | 59 | 0.198 | 1 | 0.802 | 48.12
1 |  | 1.0001 | 5.4e-05 | 1 | 27 | 1 | 2.7e-05 | D | 0.99986 |  | 45 | 44.9963 | -0.0037 | 44 | 59 | 46.68 | 161986.68 | 0.9963 | 44 | 59.778 | 59 | 0.222 | 1 | 0.778 | 46.68
1.5 |  | 1.5001 | 8.4e-05 | 1.5 | 24 | 1.5 | 2.4e-05 | E | 1.2498 |  | 60 | 59.9962 | -0.0038 | 59 | 59 | 46.32 | 215986.32 | 0.9962 | 59 | 59.772 | 59 | 0.228 | 1 | 0.772 | 46.32
2 |  | 2.0001 | 7.2e-05 | 2 | 42 | 2 | 4.2e-05 | F | 0.06239 |  | 75 | 74.9969 | -0.0031 | 74 | 59 | 48.84 | 269988.84 | 0.9969 | 74 | 59.814 | 59 | 0.186 | 1 | 0.814 | 48.84
2.5 |  | 2.5001 | 0.000116 | 2.5 | 67 | 2.5001 | 6.7e-05 | G | 0.18742 |  | 90 | 89.9996 | -0.0004 | 89 | 59 | 58.56 | 323998.56 | 0.9996 | 89 | 59.976 | 59 | 0.024 | 1 | 0.976 | 58.56
3 |  | 3.0002 | 0.000158 | 3 | 64 | 3.0001 | 6.4e-05 | H | 0.24991 |  | 105 | 104.997 | -0.003 | 104 | 59 | 49.2 | 377989.2 | 0.997 | 104 | 59.82 | 59 | 0.18 | 1 | 0.82 | 49.2
3.5 |  | 3.5002 | 0.000156 | 3.5 | 53 | 3.5001 | 5.3e-05 | J | 0.43741 |  | 120 | 119.9963 | -0.0037 | 119 | 59 | 46.68 | 431986.68 | 0.9963 | 119 | 59.778 | 59 | 0.222 | 1 | 0.778 | 46.68
4 |  | 4.0002 | 0.000191 | 4 | 52 | 4.0001 | 5.2e-05 | K | 0.6249 |  | 135 | 134.9975 | -0.0025 | 134 | 59 | 51 | 485991 | 0.9975 | 134 | 59.85 | 59 | 0.15 | 1 | 0.85 | 51
4.5 |  | 4.5002 | 0.000215 | 4.5 | 42 | 4.5 | 4.2e-05 | L | 0.06245 |  | 150 | 149.9988 | -0.0012 | 149 | 59 | 55.68 | 539995.68 | 0.9988 | 149 | 59.928 | 59 | 0.072 | 1 | 0.928 | 55.68
5 |  | 5.0002 | 0.000194 | 5 | 64 | 5.0001 | 6.4e-05 | M | 0.12495 |  | 165 | 164.9941 | -0.0059 | 164 | 59 | 38.76 | 593978.76 | 0.9941 | 164 | 59.646 | 59 | 0.354 | 1 | 0.646 | 38.76
5.5 |  | 5.5002 | 0.000213 | 5.5 | 18 | 5.5 | 1.8e-05 | N | 0.18748 |  | 180 | 179.9966 | -0.0034 | 179 | 59 | 47.76 | 647987.76 | 0.9966 | 179 | 59.796 | 59 | 0.204 | 1 | 0.796 | 47.76
6 |  | 6.0002 | 0.000231 | 6 | 27 | 6 | 2.7e-05 | O | 0.24996
6.5 |  | 6.5002 | 0.000232 |  |  |  |  | P | 0.06237
7 |  | 7.0002 | 0.000196 |  |  |  |  | Q | 0.1874
7.5 |  | 7.5002 | 0.000229 |  |  |  |  | R | 0.24988
8 |  | 8.0002 | 0.000247 |  |  | 1 |  | S | 0.3749
8.5 |  | 8.5002 | 0.000239 |  |  | 0.99998 |  | T | 0.06248
9 |  | 9.0003 | 0.000251 |  |  | -3.8e-05 |  | U | 0.12494
9.5 |  | 9.5003 | 0.000329 |  |  |  |  | V | 0.18753
10 |  | 10.0003 | 0.000327 |  |  |  |  | W | 0.0625
10.5 |  | 10.5003 | 0.000326 |  |  |  |  | Z | 0.12497
11 |  | 11.0003 | 0.000338
11.5 |  | 11.5004 | 0.000375
12 |  | 12.0004 | 0.000372
Y Scale Interp
|  |  |  | known Col. Lookup (The value from the known column for which you want a corresponding unknown-other column value) | Unknown. Column. result | Ct of numbers< known column lookup in the known column. | known column low | known column high | unknown col. Low | unknown column high
|  |  |  | 0.1 | 0.1 | 1 | 0 | 0.5 | 0 | 0.50002
|  |  |  | 0.5 | 0.50002 | 1 | 0 | 0.5 | 0 | 0.50002
|  |  |  | 1 | 1 | 2 | 0.5 | 1 | 0.50002 | 1
|  |  |  | 1.5 | 1.5 | 3 | 1 | 1.5 | 1 | 1.5
|  |  |  | 1.7 | 1.7 | 4 | 1.5 | 2 | 1.5 | 2
|  |  |  | 0 | #VALUE! | 0 | Nominal | 0 | Actual | 0
|  |  |  | 0 | #VALUE! | 0 | Nominal | 0 | Actual | 0
|  |  |  | 0 | #VALUE! | 0 | Nominal | 0 | Actual | 0
|  |  |  | 0 | #VALUE! | 0 | Nominal | 0 | Actual | 0
|  |  |  | 0 | #VALUE! | 0 | Nominal | 0 | Actual | 0
|  |  |  | 0 | #VALUE! | 0 | Nominal | 0 | Actual | 0
|  |  |  | 0 | #VALUE! | 0 | Nominal | 0 | Actual | 0
|  |  |  | 0 | #VALUE! | 0 | Nominal | 0 | Actual | 0
|  |  |  | 0 | #VALUE! | 0 | Nominal | 0 | Actual | 0
[5 empty] | 0 | 0 | 0 | 6 | 0 | 6

## OC4000 Values
[7 empty] | Due Date |  | 2027-11-22
|  |  | GAGE LINE RETICLE CERTIFICATION VALUES |  |  |  | Calibration Date |  | 2024-11-22
OC4000 VALUES
Top Inch Scale |  |  |  | Vertical Inch Scale |  |  |  | CIRCLES | ACTUAL SIZE |  | ANGLES |  |  | Standard Counts |  |  |  | minutes extracted | Extraction |  |  |  |  | seconds extracted
Nominal | error (µ") | Actual |  | Nominal | error (µ") | Actual |  | A | 0.06222 |  | Nominal | ACTUAL VALUES | Dev. | Deg | Min | Sec | Std Counts |  | Deg | Min | seconds extractor |  |  |  | Sec
0 |  | 0 | 0 | 0 |  | 0 | 0 | B | 0.2497 |  | 15 | 14.998 | -0.002 | 14 | 59 | 52.8 | 53992.8 | 0.998 | 14 | 59.88 | 59 | 0.12 | 1 | 0.88 | 52.8
0.5 |  | 0.50006 | 5.6e-05 | 0.5 | 130 | 0.50013 | 0.00013 | C | 0.6247 |  | 30 | 30.024 | 0.024 | 30 | 1 | 26.4 | 108086.4 | 0.024 | 30 | 1.44 | 1 | 58.56 | 59 | 0.44 | 26.4
1 |  | 1.0001 | 7.1e-05 | 1 | 20 | 1 | 2e-05 | D | 0.99969 |  | 45 | 45.0065 | 0.0065 | 45 | 0 | 23.4 | 162023.4 | 0.0065 | 45 | 0.39 | 0 | 59.61 | 60 | 0.39 | 23.4
1.5 |  | 1.5001 | 0.000129 | 1.5 | 200 | 1.5002 | 0.0002 | E | 1.2497 |  | 60 | 59.9925 | -0.0075 | 59 | 59 | 33 | 215973 | 0.9925 | 59 | 59.55 | 59 | 0.45 | 1 | 0.55 | 33
2 |  | 2.0002 | 0.000168 | 1.7 | 140 | 1.7001 | 0.00014 | F | 0.0622 |  | 75 | 74.999 | -0.001 | 74 | 59 | 56.4 | 269996.4 | 0.999 | 74 | 59.94 | 59 | 0.06 | 1 | 0.94 | 56.4
2.5 |  | 2.5002 | 0.000168 |  | 0 |  | 0 | G | 0.18719 |  | 90 | 89.9943 | -0.0057 | 89 | 59 | 39.48 | 323979.48 | 0.9943 | 89 | 59.658 | 59 | 0.342 | 1 | 0.658 | 39.48
3 |  | 3.0002 | 0.000206 |  | 0 |  | 0 | H | 0.2469 |  | 105 | 104.9965 | -0.0035 | 104 | 59 | 47.4 | 377987.4 | 0.9965 | 104 | 59.79 | 59 | 0.21 | 1 | 0.79 | 47.4
3.5 |  | 3.5003 | 0.000261 |  | 0 |  | 0 | J | 0.4372 |  | 120 | 120.0084 | 0.0084 | 120 | 0 | 30.24 | 432030.24 | 0.0084 | 120 | 0.504 | 0 | 59.496 | 60 | 0.504 | 30.24
4 |  | 4.0003 | 0.000273 |  | 0 |  | 0 | K | 0.62471 |  | 135 | 134.9934 | -0.0066 | 134 | 59 | 36.24 | 485976.24 | 0.9934 | 134 | 59.604 | 59 | 0.396 | 1 | 0.604 | 36.24
4.5 |  | 4.5003 | 0.000303 |  | 0 |  | 0 | L | 0.06226 |  | 150 | 150.0028 | 0.0028 | 150 | 0 | 10.08 | 540010.08 | 0.0028 | 150 | 0.168 | 0 | 59.832 | 60 | 0.168 | 10.08
5 |  | 5.0003 | 0.000301 |  | 0 |  | 0 | M | 0.12474 |  | 165 | 165.0062 | 0.0062 | 165 | 0 | 22.32 | 594022.32 | 0.0062 | 165 | 0.372 | 0 | 59.628 | 60 | 0.372 | 22.32
5.5 |  | 5.5003 | 0.000326 |  | 0 |  | 0 | N | 0.18724 |  | 180 | 179.9951 | -0.0049 | 179 | 59 | 42.36 | 647982.36 | 0.9951 | 179 | 59.706 | 59 | 0.294 | 1 | 0.706 | 42.36
6 |  | 6.0003 | 0.000349 |  | 0 |  | 0 | O | 0.24974
6.5 |  | 6.5003 | 0.000349 |  |  |  |  | P | 0.06227
7 |  | 7.0004 | 0.000361 |  |  |  |  | Q | 0.18724
7.5 |  | 7.5004 | 0.000377 |  |  |  |  | R | 0.24974
8 |  | 8.0004 | 0.000386 |  |  | -1.7001 |  | S | 0.37475
8.5 |  | 8.5004 | 0.000415 |  |  | 0.99998 |  | T | 0.06226
9 |  | 9.0004 | 0.000439 |  |  | 2.7001 |  | U | 0.12474
9.5 |  | 9.5005 | 0.000481 |  |  |  |  | V | 0.18724
10 |  | 10.0005 | 0.000477 |  |  |  |  | W | 0.06232
10.5 |  | 10.5005 | 0.000535 |  |  |  |  | Z | 0.12483
11 |  | 11.0005 | 0.000537
11.5 |  | 11.5006 | 0.000573
12 |  | 12.0006 | 0.000619
Y Scale Interp
|  |  |  | known Col. Lookup (The value from the known column for which you want a corresponding unknown-other column value) | Unknown. Column. result | Ct of numbers< known column lookup in the known column. | known column low | known column high | unknown col. Low | unknown column high
|  |  |  | 0.1 | 0.10003 | 1 | 0 | 0.5 | 0 | 0.50013
|  |  |  | 0.5 | 0.50013 | 1 | 0 | 0.5 | 0 | 0.50013
|  |  |  | 1 | 1 | 2 | 0.5 | 1 | 0.50013 | 1
|  |  |  | 1.5 | 1.5002 | 3 | 1 | 1.5 | 1 | 1.5002
|  |  |  | 1.7 | 1.7001 | 4 | 1.5 | 1.7 | 1.5002 | 1.7001
|  |  |  | 0 | #VALUE! | 0 | Nominal | 0 | Actual | 0
|  |  |  | 0 | #VALUE! | 0 | Nominal | 0 | Actual | 0
|  |  |  | 0 | #VALUE! | 0 | Nominal | 0 | Actual | 0
|  |  |  | 0 | #VALUE! | 0 | Nominal | 0 | Actual | 0
|  |  |  | 0 | #VALUE! | 0 | Nominal | 0 | Actual | 0
|  |  |  | 0 | #VALUE! | 0 | Nominal | 0 | Actual | 0
|  |  |  | 0 | #VALUE! | 0 | Nominal | 0 | Actual | 0
|  |  |  | 0 | #VALUE! | 0 | Nominal | 0 | Actual | 0
|  |  |  | 0 | #VALUE! | 0 | Nominal | 0 | Actual | 0
[5 empty] | #DIV/0! | 0 | 0 | 0 | 0 | 0

## Vaidation Log
Date: | Action:
2015-03-11 | Lookup Corrections verified, Magnification function verified, Password protection added
2015-06-17 | OC2000 values added, programmed. Fillable cells identified by color code. Worksheet renamed "Optical Comparator Worksheet"
2015-01-09 | Magscale recalibrated, Mag accuracy deviation and tolerance calculation was corrected. Deviations calculated for angles, Visual improvements, Summary of results page added
2016-01-19 | Corrected a Y axis formula error
2017-02-07 | Removed Table aligmnent fields
2018-07-28 | Added OC3000, made start point for x and y axis adjustable
2021-07-19 | Updated reference values, no software changes
2024-06-03 | Updated reference values, updated OC3000 standard entry format
2024-07-03 | added decimal degrees deviation
2025-11-06 | fixed a calculation error in the minutes and seconds extraction of the OC3000 and OC4000 tabs (angle)

## OC1000 Values
OC1000 Cert Values
OC1000 VALUES |  |  | =
Top Inch Scale |  |  |  | Vertical Inch Scale |  |  |  | CIRCLES | ACTUAL SIZE | ANGLES |  |  | Standard Counts
Nominal | error (µ") | Actual |  | Nominal | error (µ") | Actual |  | A |  | Nominal | ACTUAL VALUES | Dev. | Deg | Min | Sec | Std Counts
0 |  | 0 | 0 | 0 |  | 0 | 0 | B | 0.25001 | 15 | 14° 59' 52" | +5" | 14 | 59 | 52 | 53992
0.5 | 63 | 0.50006 | 0 | 0.2 | 1 | 0.2 | 1e-06 | C | 0.62503 | 30 |  | -40" |  |  |  | 0
1 | 134 | 1.0001 | 0 | 0.4 | 14 | 0.40001 | 1.4e-05 | D |  | 45 |  | +9" |  |  |  | 0
1.5 | 171 | 1.5002 | 0 | 0.6 | 19 | 0.60002 | 1.9e-05 | E |  | 60 |  | +34" |  |  |  | 0
2 | 213 | 2.0002 | 0 | 0.8 | 36 | 0.80004 | 3.6e-05 | F | 0.06253 | 75 |  | -39" |  |  |  | 0
2.5 | 251 | 2.5003 |  | 1 | 31 | 1 | 3.1e-05 | G |  | 90 | 90° 0' 04" | -1' 8" | 90 | 0 | 4 | 324004
3 | 293 | 3.0003 | 0 | 1.1 | 40 | 1.1 | 4e-05 | H | 0.25003 | 105 |  | -59" |  |  |  | 0
3.5 | 319 | 3.5003 |  | 1.2 | 39 | 1.2 | 3.9e-05 | J |  | 120 |  | -49" |  |  |  | 0
4 | 339 | 4.0003 | 0 | 1.3 | 30 | 1.3 | 3e-05 | K |  | 135 | 135° 0' 25" | -11" | 135 | 0 | 25 | 486025
4.5 | 409 | 4.5004 |  | 1.4 | 39 | 1.4 | 3.9e-05 | L |  | 150 |  | -1' 21" |  |  |  | 0
5 | 425 | 5.0004 | 0 | 1.5 | 45 | 1.5 | 4.5e-05 | M |  | 165 |  | -48" |  |  |  | 0
5.5 | 458 | 5.5005 |  | 1.6 | 43 | 1.6 | 4.3e-05 | N |  | 180 |  | -26" |  |  |  | 0
6 | 449 | 6.0004 | 0 | 1.7 | 43 | 1.7 | 4.3e-05 | O
6.5 | 467 | 6.5005 | [5 empty] | P
7 | 465 | 7.0005 | 0 |  |  |  |  | Q
7.5 | 475 | 7.5005 | [5 empty] | R
8 | 513 | 8.0005 | 0 |  |  |  |  | S
8.5 | [7 empty] | T
9 |  |  | 0 |  |  |  |  | U
9.5 | [7 empty] | V
10 |  |  | 0 |  |  |  |  | W
10.5 | [7 empty] | Z
11 |  |  | 0
11.5
12 |  |  | 0

## OC1000 Calibration
Standardization of OC | [12 empty] | OC1000 Values
Master Scale Used:
Date: | [5 empty] | A,B,C,G,J,K
X axis |  |  | Y Axis |  |  | Angles | [6 empty] | X axis |  |  | Y Axis |  |  | Angles
Nominal | indicated | True Error | Nominal | indicated | True Error | Nominal | Deg | Min | Sec | Cts | Std |  | Nominal | Indicated | True Error | Nominal | Indicated | True Error | Nominal | Deg | Min | Sec | Cts | True Rdg | Deg | Min | Sec
0.5 | 0.50001 | -1e-06 | 0.1 | 0.09995 | -0.10005 | 15 | 15 | 0 | 37 | 54037 | 53996.04 |  | 0.5 | 0.50004 | 4.1e-05 | 0.1 | 0.09999 | 0.10004 | 15 | 15 | 0 | 56 | 54056 | 54015.04
1 | 1.0001 | 0 | 0.5 | 0.49991 | -0.50011 | 30 | 30 | 0 | 32 | 108032 | 107988.12 |  | 1 | 1.0001 | 8e-05 | 0.5 | 0.49995 | 0.50006 | 30 | 30 | 0 | 30 | 108030 | 107986.12
1.5 | 1.5001 | 0 | 1 | 0.99989 | -1.0001 | 90 | 90 | 0 | 51 | 324051 | 323998.56 |  | 1.5 | 1.5001 | 5e-05 | 1 | 0.99993 | 1.0001 | 90 | 90 | 0 | 53 | 324053 | 324000.56
2 | 2.0001 | 1.5e-05 | 1.5 | 1.4999 | -1.5001 | 135 | 135 | 1 | 8 | 486068 | 485991 |  | 2 | 2.0001 | 0.000115 | 1.5 | 1.4999 | 1.5 | 135 | 135 | 2 | 4 | 486124 | 486047
2.5 | 2.5001 | 1.6e-05 | 1.7 | 1.7 | -1.7001 | 150 | 150 | 0 | 49 | 540049 | 539995.68 |  | 2.5 | 2.5001 | 7.4e-05 | 1.7 | 1.7001 | 1.7001 | 150 | 150 | 1 | 11 | 540071 | 540017.68
3 | 3.0002 | 7e-06 |  |  | #VALUE! | [7 empty] | 3 | 3.0002 | 0.000163 | 0 |  | #VALUE! | 0 |  |  |  |  | 0
3.5 | 3.5002 | 2e-05 |  |  | #VALUE! | [7 empty] | 3.5 | 3.5002 | 0.00016 | 0 |  | #VALUE! | 0 |  |  |  |  | 0
4 | 4.0002 | 5e-06 |  |  | #VALUE! | [7 empty] | 4 | 4.0002 | 0.000215 | 0 |  | #VALUE! | 0 |  |  |  |  | 0
4.5 | 4.5002 | 1.7e-05 |  |  | #VALUE! | [7 empty] | 4.5 | 4.5003 | 0.000283 | 0 |  | #VALUE! | 0 |  |  |  |  | 0
5 | 5.0002 | 4.1e-05 |  |  | #VALUE! | [7 empty] | 5 | 5.0003 | 0.000259 | 0 |  | #VALUE! | 0 |  |  |  |  | 0
5.5 | 5.5002 | 2.1e-05 |  |  | #VALUE! | [7 empty] | 5.5 | 5.5004 | 0.000329 | 0 |  | #VALUE! | 0 |  |  |  |  | 0
6 | 6.0003 | 3.9e-05 |  |  | #VALUE! | [7 empty] | 6 | 6.0004 | 0.000391 | 0 |  | #VALUE! | 0 |  |  |  |  | 0
6.5 | 6.5003 | 3.8e-05 |  |  | #VALUE! | [7 empty] | 6.5 | 6.5004 | 0.000392 | 0 |  | #VALUE!
7 | 7.0002 | 4.8e-05 |  |  | #VALUE! |  |  | Mag Circles |  |  |  |  | 7 | 7.0004 | 0.000382 | 0 |  | #VALUE! |  |  | Mag Circles
7.5 | 7.5003 | 4.7e-05 |  |  |  |  | Circle | Nominal | indicated | True Error |  |  | 7.5 | 7.5005 | 0.000433 |  |  |  |  | Circle | Nominal | Indicated | True Error
8 | 8.0003 | 1.7e-05 |  |  |  |  | A |  |  |  |  | reset at 7 | 8 | 8.0005 | 0.000503 |  |  |  |  | A
8.5 | [6 empty] | B |  |  |  |  | 7.5005 | 8.5 | 8.5005 | 0.000433 |  |  |  |  | B
9 | [6 empty] | C |  |  |  |  | 8.0005 | 9 | 9.0005 | 0.000463 |  |  |  |  | C
9.5 | [6 empty] | D | [12 empty] | D
10 | [6 empty] | E | [12 empty] | E
10.5 | [6 empty] | F | [12 empty] | F
11 | [6 empty] | G | [12 empty] | G
11.5 | [6 empty] | H | [12 empty] | H
12 | [6 empty] | I | [12 empty] | I
[7 empty] | J | [12 empty] | J
[7 empty] | K | [12 empty] | K
[7 empty] | L | [12 empty] | L
[7 empty] | M | [12 empty] | M
[7 empty] | N | [12 empty] | N
[7 empty] | O | [12 empty] | O

## OC100A Results 2019
Magnification Scale Calibration Values
| 10 | X setting | Cal. Date: | 2016-01-08 |  |  | Interpolated correction factors for Normal Use
[7 empty] | known Col. Lookup (The value from the known column for which you want a corresponding unknown-other column value) | Unknown. Column. result | Ct of numbers< known column lookup in the known column. | known column low | known column high | unknown col. Low | unknown column high | Actual Result | Comparator Characterization
| Marked | Nominal | indicated | Error | Cal. Corr. Factor | Usage Corr. Factor | 0 | #VALUE! | 0 | Marked | 0 | Usage Corr. Factor | 0 | #VALUE! | Nominal | Indicated | Correction from OC2000 | Corrected Value | Characterized Correction
| 0 | 0 | 0 | 0 |  |  | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 0 | 0 | 0 | 0
| 0.02 | 0.2 |  | -0.2 | -1.04e-05 | 0.20001 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 0.5 | 0.50004 | -1.4e-05 | 0.50003 | -2.6e-05
| 0.04 | 0.4 |  | -0.4 | -2.08e-05 | 0.40002 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 1 | 1 | -5.4e-05 | 0.99996 | 4.4e-05
| 0.06 | 0.6 |  | -0.6 | -1.2e-05 | 0.60001 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 1.5 | 1.5001 | -8.4e-05 | 1.5 | -4.6e-05
| 0.08 | 0.8 |  | -0.8 | 1.6e-05 | 0.79998 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 2 | 2.0001 | -7.2e-05 | 2 | -1.8e-05
| 0.1 | 1 | 0.99989 | -0.00011 | 4.4e-05 | 6.6e-05 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 2.5 | 2.5002 | -0.000116 | 2.5001 | -5.4e-05
| 0.12 | 1.2 |  | -1.2 | 8e-06 | 1.2 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 3 | 3.0002 | -0.000158 | 3.0001 | -9.2e-05
| 0.14 | 1.4 |  | -1.4 | -2.8e-05 | 1.4 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 3.5 | 3.5003 | -0.000156 | 3.5001 | -0.000144
| 0.16 | 1.6 |  | -1.6 | -4.04e-05 | 1.6 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 4 | 4.0004 | -0.000191 | 4.0002 | -0.000189
| 0.18 | 1.8 |  | -1.8 | -2.92e-05 | 1.8 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 4.5 | 4.5004 | -0.000215 | 4.5002 | -0.000165
| 0.2 | 2 | 1.9999 | -7e-05 | -1.8e-05 | 8.8e-05 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 5 | 5.0004 | -0.000194 | 5.0002 | -0.000226
| 0.22 | 2.2 |  | -2.2 | -3.24e-05 | 2.2 | [8 empty] | 5.5 | 5.5005 | -0.000213 | 5.5002 | -0.000247
| 0.24 | 2.4 |  | -2.4 | -4.68e-05 | 2.4 | Interp. Correction factors for obtaining TRUE OC100 Error | [7 empty] | 6 | 6.0005 | -0.000231 | 6.0002 | -0.000239
| 0.26 | 2.6 |  | -2.6 | -6.16e-05 | 2.6001 | known Col. Lookup (The value from the known column for which you want a corresponding unknown-other column value) | Unknown. Column. result | Ct of numbers< known column lookup in the known column. | known column low | known column high | unknown col. Low | unknown column high |  | 6.5 | 6.5004 | -0.000232 | 6.5002 | -0.000198
| 0.28 | 2.8 |  | -2.8 | -7.68e-05 | 2.8001 | 0.2 | -1.04e-05 | 1 | 0 | 0.5 | 0 | -2.6e-05 |  | 7 | 7.0005 | -0.000196 | 7.0004 | -0.000354
| 0.3 | 3 | 3.0001 | 0.0001 | -9.2e-05 | -8e-06 | 0.4 | -2.08e-05 | 1 | 0 | 0.5 | 0 | -2.6e-05 |  | 7.5 | 7.5006 | -0.000229 | 7.5004 | -0.000361
| 0.32 | 3.2 |  | -3.2 | -0.0001128 | 3.2001 | 0.6 | -1.2e-05 | 2 | 0.5 | 1 | -2.6e-05 | 4.4e-05 |  | 8 | 8.0006 | -0.000247 | 8.0004 | -0.000353
| 0.34 | 3.4 |  | -3.4 | -0.0001336 | 3.4001 | 0.8 | 1.6e-05 | 2 | 0.5 | 1 | -2.6e-05 | 4.4e-05
| 0.36 | 3.6 |  | -3.6 | -0.000153 | 3.6002 | 1 | 4.4e-05 | 2 | 0.5 | 1 | -2.6e-05 | 4.4e-05
| 0.38 | 3.8 |  | -3.8 | -0.000171 | 3.8002 | 1.2 | 8e-06 | 3 | 1 | 1.5 | 4.4e-05 | -4.6e-05
| 0.4 | 4 | 4.0002 | 0.00022 | -0.000189 | -3.1e-05 | 1.4 | -2.8e-05 | 3 | 1 | 1.5 | 4.4e-05 | -4.6e-05
| 0.42 | 4.2 |  | -4.2 | -0.0001794 | 4.2002 | 1.6 | -4.04e-05 | 4 | 1.5 | 2 | -4.6e-05 | -1.8e-05
| 0.44 | 4.4 |  | -4.4 | -0.0001698 | 4.4002 | 1.8 | -2.92e-05 | 4 | 1.5 | 2 | -4.6e-05 | -1.8e-05
| 0.46 | 4.6 |  | -4.6 | -0.0001772 | 4.6002 | 2 | -1.8e-05 | 4 | 1.5 | 2 | -4.6e-05 | -1.8e-05
| 0.48 | 4.8 |  | -4.8 | -0.0002016 | 4.8002 | 2.2 | -3.24e-05 | 5 | 2 | 2.5 | -1.8e-05 | -5.4e-05
| 0.5 | 5 | 5.0006 | 0.00064 | -0.000226 | -0.000414 | 2.4 | -4.68e-05 | 5 | 2 | 2.5 | -1.8e-05 | -5.4e-05
| 0.52 | 5.2 |  | -5.2 | -0.0002344 | 5.2002 | 2.6 | -6.16e-05 | 6 | 2.5 | 3 | -5.4e-05 | -9.2e-05
| 0.54 | 5.4 |  | -5.4 | -0.0002428 | 5.4002 | 2.8 | -7.68e-05 | 6 | 2.5 | 3 | -5.4e-05 | -9.2e-05
| 0.56 | 5.6 |  | -5.6 | -0.0002454 | 5.6002 | 3 | -9.2e-05 | 6 | 2.5 | 3 | -5.4e-05 | -9.2e-05
| 0.58 | 5.8 |  | -5.8 | -0.0002422 | 5.8002 | 3.2 | -0.0001128 | 7 | 3 | 3.5 | -9.2e-05 | -0.000144
| 0.6 | 6 | 6.0002 | 0.00019 | -0.000239 | 4.9e-05 | 3.4 | -0.0001336 | 7 | 3 | 3.5 | -9.2e-05 | -0.000144
| 0.62 | 6.2 |  | -6.2 | -0.0002226 | 6.2002 | 3.6 | -0.000153 | 8 | 3.5 | 4 | -0.000144 | -0.000189
| 0.64 | 6.4 |  | -6.4 | -0.0002062 | 6.4002 | 3.8 | -0.000171 | 8 | 3.5 | 4 | -0.000144 | -0.000189
| 0.66 | 6.6 |  | -6.6 | -0.0002292 | 6.6002 | 4 | -0.000189 | 8 | 3.5 | 4 | -0.000144 | -0.000189
| 0.68 | 6.8 |  | -6.8 | -0.0002916 | 6.8003 | 4.2 | -0.0001794 | 9 | 4 | 4.5 | -0.000189 | -0.000165
| 0.7 | 7 | 7.0004 | 0.00043 | -0.000354 | -7.6e-05 | 4.4 | -0.0001698 | 9 | 4 | 4.5 | -0.000189 | -0.000165
| 0.72 | 7.2 |  | -7.2 | -0.0003568 | 7.2004 | 4.6 | -0.0001772 | 10 | 4.5 | 5 | -0.000165 | -0.000226
| 0.74 | 7.4 |  | -7.4 | -0.0003596 | 7.4004 | 4.8 | -0.0002016 | 10 | 4.5 | 5 | -0.000165 | -0.000226
| 0.76 | 7.6 |  | -7.6 | -0.0003594 | 7.6004 | 5 | -0.000226 | 10 | 4.5 | 5 | -0.000165 | -0.000226
| 0.78 | 7.8 |  | -7.8 | -0.0003562 | 7.8004 | 5.2 | -0.0002344 | 11 | 5 | 5.5 | -0.000226 | -0.000247
| 0.8 | 8 | 8.0007 | 0.00068 | -0.000353 | -0.000327 | 5.4 | -0.0002428 | 11 | 5 | 5.5 | -0.000226 | -0.000247
| 0.82 | 8.2 |  | -8.2 | -0.00036183 | 8.2 | 5.6 | -0.0002454 | 12 | 5.5 | 6 | -0.000247 | -0.000239
| 0.84 | 8.4 |  | -8.4 | -0.00037065 | 8.4 | 5.8 | -0.0002422 | 12 | 5.5 | 6 | -0.000247 | -0.000239
| 0.86 | 8.6 |  | -8.6 | -0.00037948 | 8.6 | 6 | -0.000239 | 12 | 5.5 | 6 | -0.000247 | -0.000239
| 0.88 | 8.8 |  | -8.8 | -0.0003883 | 8.8 | 6.2 | -0.0002226 | 13 | 6 | 6.5 | -0.000239 | -0.000198
0.99989 | 0.9 | 9 | 9.0006 | 0.00057 | -0.00039713 | -0.000614 | 6.4 | -0.0002062 | 13 | 6 | 6.5 | -0.000239 | -0.000198
| 0.92 | 9.2 |  | -9.2 | -0.00040595 | 9.2 | 6.6 | -0.0002292 | 14 | 6.5 | 7 | -0.000198 | -0.000354
| 0.94 | 9.4 |  | -9.4 | -0.00041478 | 9.4 | 6.8 | -0.0002916 | 14 | 6.5 | 7 | -0.000198 | -0.000354
| 0.96 | 9.6 |  | -9.6 | -0.0004236 | 9.6 | 7 | -0.000354 | 14 | 6.5 | 7 | -0.000198 | -0.000354
| 0.98 | 9.8 |  | -9.8 | -0.00043243 | 9.8 | 7.2 | -0.0003568 | 15 | 7 | 7.5 | -0.000354 | -0.000361
2.0002 | 1 | 10 | 10.0009 | 0.00089 | -0.00044125 | -0.000872 | 7.4 | -0.0003596 | 15 | 7 | 7.5 | -0.000354 | -0.000361
| 1.02 | 10.2 |  | -10.2 | -0.00045008 | 10.2 | 7.6 | -0.0003594 | 16 | 7.5 | 8 | -0.000361 | -0.000353
| 1.04 | 10.4 |  | -10.4 | -0.0004589 | 10.4 | 7.8 | -0.0003562 | 16 | 7.5 | 8 | -0.000361 | -0.000353
| 1.06 | 10.6 |  | -10.6 | -0.00046773 | 10.6001 | 8 | -0.000353 | 16 | 7.5 | 8 | -0.000361 | -0.000353
| 1.08 | 10.8 |  | -10.8 | -0.00047655 | 10.8001 | 8.2 | -0.00036183 | 17 | 8 | 0 | -0.000353 | 0
2.9998 | 1.1 | 11 | 11.0005 | 0.00046 | -0.00048538 | -0.000368 | 8.4 | -0.00037065 | 17 | 8 | 0 | -0.000353 | 0
| 1.12 | 11.2 |  | -11.2 | -0.0004942 | 11.2001 | 8.6 | -0.00037948 | 17 | 8 | 0 | -0.000353 | 0
| 1.14 | 11.4 |  | -11.4 | -0.00050303 | 11.4001 | 8.8 | -0.0003883 | 17 | 8 | 0 | -0.000353 | 0
| 1.16 | 11.6 |  | -11.6 | -0.00051185 | 11.6002 | 9 | -0.00039713 | 17 | 8 | 0 | -0.000353 | 0
| 1.18 | 11.8 |  | -11.8 | -0.00052068 | 11.8002 | 9.2 | -0.00040595 | 17 | 8 | 0 | -0.000353 | 0
4.0002 | 1.2 | 12 | 12.0009 | 0.0009 | -0.0005295 | -0.000711 | 9.4 | -0.00041478 | 17 | 8 | 0 | -0.000353 | 0
[7 empty] | 9.6 | -0.0004236 | 17 | 8 | 0 | -0.000353 | 0
[7 empty] | 9.8 | -0.00043243 | 17 | 8 | 0 | -0.000353 | 0
[7 empty] | 10 | -0.00044125 | 17 | 8 | 0 | -0.000353 | 0
[7 empty] | 10.2 | -0.00045008 | 17 | 8 | 0 | -0.000353 | 0
[7 empty] | 10.4 | -0.0004589 | 17 | 8 | 0 | -0.000353 | 0
[7 empty] | 10.6 | -0.00046773 | 17 | 8 | 0 | -0.000353 | 0
[7 empty] | 10.8 | -0.00047655 | 17 | 8 | 0 | -0.000353 | 0
[7 empty] | 11 | -0.00048538 | 17 | 8 | 0 | -0.000353 | 0
[7 empty] | 11.2 | -0.0004942 | 17 | 8 | 0 | -0.000353 | 0
[7 empty] | 11.4 | -0.00050303 | 17 | 8 | 0 | -0.000353 | 0
[7 empty] | 11.6 | -0.00051185 | 17 | 8 | 0 | -0.000353 | 0
[7 empty] | 11.8 | -0.00052068 | 17 | 8 | 0 | -0.000353 | 0
[7 empty] | 12 | -0.0005295 | 17 | 8 | 0 | -0.000353 | 0
[7 empty] | 0 | #VALUE! | 0 | Nominal | 0 | Characterized Correction | 0
[7 empty] | 0 | #VALUE! | 0 | Nominal | 0 | Characterized Correction | 0

## Magscale OC100 Old test
Magnification Scale Calibration Values
10 | X setting | Cal. Date: | 2016-01-08 |  |  | Interpolated correction factors for Normal Use
[6 empty] | known Col. Lookup (The value from the known column for which you want a corresponding unknown-other column value) | Unknown. Column. result | Ct of numbers< known column lookup in the known column. | known column low | known column high | unknown col. Low | unknown column high | Actual Result | Comparator Characterization
Marked | Nominal | indicated | Error | Cal. Corr. Factor | Usage Corr. Factor | 0 | #VALUE! | 0 | Marked | 0 | Usage Corr. Factor | 0 | #VALUE! | Nominal | Indicated | Correction from OC2000 | Corrected Value | Characterized Correction
0 | 0 | 0 | 0 |  |  | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 0 | 0 | 0 | 0
0.02 | 0.2 | 0.20014 | 0.00014 | -1.04e-05 | -0.0001296 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 0.5 | 0.50004 | -1.4e-05 | 0.50003 | -2.6e-05
0.04 | 0.4 | 0.40023 | 0.00023 | -2.08e-05 | -0.0002092 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 1 | 1 | -5.4e-05 | 0.99996 | 4.4e-05
0.06 | 0.6 | 0.60021 | 0.00021 | -1.2e-05 | -0.000198 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 1.5 | 1.5001 | -8.4e-05 | 1.5 | -4.6e-05
0.08 | 0.8 | 0.80034 | 0.00034 | 1.6e-05 | -0.000356 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 2 | 2.0001 | -7.2e-05 | 2 | -1.8e-05
0.1 | 1 | 1.0001 | 5e-05 | 4.4e-05 | -9.4e-05 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 2.5 | 2.5002 | -0.000116 | 2.5001 | -5.4e-05
0.12 | 1.2 | 1.1999 | -0.0001 | 8e-06 | 9.2e-05 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 3 | 3.0002 | -0.000158 | 3.0001 | -9.2e-05
0.14 | 1.4 | 1.4001 | 8e-05 | -2.8e-05 | -5.2e-05 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 3.5 | 3.5003 | -0.000156 | 3.5001 | -0.000144
0.16 | 1.6 | 1.5999 | -6e-05 | -4.04e-05 | 0.0001004 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 4 | 4.0004 | -0.000191 | 4.0002 | -0.000189
0.18 | 1.8 | 1.8002 | 0.00019 | -2.92e-05 | -0.0001608 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 4.5 | 4.5004 | -0.000215 | 4.5002 | -0.000165
0.2 | 2 | 2.0002 | 0.00025 | -1.8e-05 | -0.000232 | 0 | #VALUE! | 0 | Marked | 0 | indicated | 0 |  | 5 | 5.0004 | -0.000194 | 5.0002 | -0.000226
0.22 | 2.2 | 2.2003 | 0.00026 | -3.24e-05 | -0.0002276 | [8 empty] | 5.5 | 5.5005 | -0.000213 | 5.5002 | -0.000247
0.24 | 2.4 | 2.4002 | 0.00024 | -4.68e-05 | -0.0001932 | Interp. Correction factors for obtaining TRUE OC100 Error | [7 empty] | 6 | 6.0005 | -0.000231 | 6.0002 | -0.000239
0.26 | 2.6 | 2.6002 | 0.00022 | -6.16e-05 | -0.0001584 | known Col. Lookup (The value from the known column for which you want a corresponding unknown-other column value) | Unknown. Column. result | Ct of numbers< known column lookup in the known column. | known column low | known column high | unknown col. Low | unknown column high |  | 6.5 | 6.5004 | -0.000232 | 6.5002 | -0.000198
0.28 | 2.8 | 2.8003 | 0.00028 | -7.68e-05 | -0.0002032 | 0.2 | -1.04e-05 | 1 | 0 | 0.5 | 0 | -2.6e-05 |  | 7 | 7.0005 | -0.000196 | 7.0004 | -0.000354
0.3 | 3 | 3.0002 | 0.00021 | -9.2e-05 | -0.000118 | 0.4 | -2.08e-05 | 1 | 0 | 0.5 | 0 | -2.6e-05 |  | 7.5 | 7.5006 | -0.000229 | 7.5004 | -0.000361
0.32 | 3.2 | 3.2004 | 0.00039 | -0.0001128 | -0.0002772 | 0.6 | -1.2e-05 | 2 | 0.5 | 1 | -2.6e-05 | 4.4e-05 |  | 8 | 8.0006 | -0.000247 | 8.0004 | -0.000353
0.34 | 3.4 | 3.4005 | 0.00048 | -0.0001336 | -0.0003464 | 0.8 | 1.6e-05 | 2 | 0.5 | 1 | -2.6e-05 | 4.4e-05
0.36 | 3.6 | 3.6004 | 0.00038 | -0.000153 | -0.000227 | 1 | 4.4e-05 | 2 | 0.5 | 1 | -2.6e-05 | 4.4e-05
0.38 | 3.8 | 3.8007 | 0.00068 | -0.000171 | -0.000509 | 1.2 | 8e-06 | 3 | 1 | 1.5 | 4.4e-05 | -4.6e-05
0.4 | 4 | 4.0005 | 0.0005 | -0.000189 | -0.000311 | 1.4 | -2.8e-05 | 3 | 1 | 1.5 | 4.4e-05 | -4.6e-05
0.42 | 4.2 | 4.2006 | 0.00059 | -0.0001794 | -0.0004106 | 1.6 | -4.04e-05 | 4 | 1.5 | 2 | -4.6e-05 | -1.8e-05
0.44 | 4.4 | 4.4006 | 0.00065 | -0.0001698 | -0.0004802 | 1.8 | -2.92e-05 | 4 | 1.5 | 2 | -4.6e-05 | -1.8e-05
0.46 | 4.6 | 4.6006 | 0.00062 | -0.0001772 | -0.0004428 | 2 | -1.8e-05 | 4 | 1.5 | 2 | -4.6e-05 | -1.8e-05
0.48 | 4.8 | 4.8007 | 0.00068 | -0.0002016 | -0.0004784 | 2.2 | -3.24e-05 | 5 | 2 | 2.5 | -1.8e-05 | -5.4e-05
0.5 | 5 | 5.0009 | 0.00086 | -0.000226 | -0.000634 | 2.4 | -4.68e-05 | 5 | 2 | 2.5 | -1.8e-05 | -5.4e-05
0.52 | 5.2 | 5.2009 | 0.00087 | -0.0002344 | -0.0006356 | 2.6 | -6.16e-05 | 6 | 2.5 | 3 | -5.4e-05 | -9.2e-05
0.54 | 5.4 | 5.4009 | 0.00093 | -0.0002428 | -0.0006872 | 2.8 | -7.68e-05 | 6 | 2.5 | 3 | -5.4e-05 | -9.2e-05
0.56 | 5.6 | 5.6009 | 0.00091 | -0.0002454 | -0.0006646 | 3 | -9.2e-05 | 6 | 2.5 | 3 | -5.4e-05 | -9.2e-05
0.58 | 5.8 | 5.8011 | 0.00112 | -0.0002422 | -0.0008778 | 3.2 | -0.0001128 | 7 | 3 | 3.5 | -9.2e-05 | -0.000144
0.6 | 6 | 6.0007 | 0.00074 | -0.000239 | -0.000501 | 3.4 | -0.0001336 | 7 | 3 | 3.5 | -9.2e-05 | -0.000144
0.62 | 6.2 | 6.2008 | 0.00076 | -0.0002226 | -0.0005374 | 3.6 | -0.000153 | 8 | 3.5 | 4 | -0.000144 | -0.000189
0.64 | 6.4 | 6.4008 | 0.00082 | -0.0002062 | -0.0006138 | 3.8 | -0.000171 | 8 | 3.5 | 4 | -0.000144 | -0.000189
0.66 | 6.6 | 6.6009 | 0.00087 | -0.0002292 | -0.0006408 | 4 | -0.000189 | 8 | 3.5 | 4 | -0.000144 | -0.000189
0.68 | 6.8 | 6.801 | 0.00097 | -0.0002916 | -0.0006784 | 4.2 | -0.0001794 | 9 | 4 | 4.5 | -0.000189 | -0.000165
0.7 | 7 | 7.0009 | 0.00095 | -0.000354 | -0.000596 | 4.4 | -0.0001698 | 9 | 4 | 4.5 | -0.000189 | -0.000165
0.72 | 7.2 | 7.2012 | 0.0012 | -0.0003568 | -0.0008432 | 4.6 | -0.0001772 | 10 | 4.5 | 5 | -0.000165 | -0.000226
0.74 | 7.4 | 7.4011 | 0.00114 | -0.0003596 | -0.0007804 | 4.8 | -0.0002016 | 10 | 4.5 | 5 | -0.000165 | -0.000226
0.76 | 7.6 | 7.6011 | 0.00108 | -0.0003594 | -0.0007206 | 5 | -0.000226 | 10 | 4.5 | 5 | -0.000165 | -0.000226
0.78 | 7.8 | 7.8013 | 0.00129 | -0.0003562 | -0.0009338 | 5.2 | -0.0002344 | 11 | 5 | 5.5 | -0.000226 | -0.000247
0.8 | 8 | 8.0011 | 0.00107 | -0.000353 | -0.000717 | 5.4 | -0.0002428 | 11 | 5 | 5.5 | -0.000226 | -0.000247
0.82 | 8.2 | 8.2002 | 0.00021 | -0.00036183 | -0.0001996 | 5.6 | -0.0002454 | 12 | 5.5 | 6 | -0.000247 | -0.000239
0.84 | 8.4 | 8.4003 | 0.00027 | -0.00037065 | -0.0002492 | 5.8 | -0.0002422 | 12 | 5.5 | 6 | -0.000247 | -0.000239
0.86 | 8.6 | 8.6003 | 0.00025 | -0.00037948 | -0.000238 | 6 | -0.000239 | 12 | 5.5 | 6 | -0.000247 | -0.000239
0.88 | 8.8 | 8.8003 | 0.0003 | -0.0003883 | -0.000316 | 6.2 | -0.0002226 | 13 | 6 | 6.5 | -0.000239 | -0.000198
0.9 | 9 | 9 | 5e-05 | -0.00039713 | -9.4e-05 | 6.4 | -0.0002062 | 13 | 6 | 6.5 | -0.000239 | -0.000198
0.92 | 9.2 | 9.2001 | 0.0001 | -0.00040595 | -0.000108 | 6.6 | -0.0002292 | 14 | 6.5 | 7 | -0.000198 | -0.000354
0.94 | 9.4 | 9.4002 | 0.00016 | -0.00041478 | -0.000132 | 6.8 | -0.0002916 | 14 | 6.5 | 7 | -0.000198 | -0.000354
0.96 | 9.6 | 9.6003 | 0.00025 | -0.0004236 | -0.0002096 | 7 | -0.000354 | 14 | 6.5 | 7 | -0.000198 | -0.000354
0.98 | 9.8 | 9.8003 | 0.00035 | -0.00043243 | -0.0003208 | 7.2 | -0.0003568 | 15 | 7 | 7.5 | -0.000354 | -0.000361
1 | 10 | 10.0004 | 0.00044 | -0.00044125 | -0.000422 | 7.4 | -0.0003596 | 15 | 7 | 7.5 | -0.000354 | -0.000361
1.02 | 10.2 | 10.2005 | 0.00054 | -0.00045008 | -0.0005076 | 7.6 | -0.0003594 | 16 | 7.5 | 8 | -0.000361 | -0.000353
1.04 | 10.4 | 10.4006 | 0.0006 | -0.0004589 | -0.0005532 | 7.8 | -0.0003562 | 16 | 7.5 | 8 | -0.000361 | -0.000353
1.06 | 10.6 | 10.6004 | 0.00038 | -0.00046773 | -0.0003184 | 8 | -0.000353 | 16 | 7.5 | 8 | -0.000361 | -0.000353
1.08 | 10.8 | 10.8005 | 0.00055 | -0.00047655 | -0.0004732 | 8.2 | -0.00036183 | 17 | 8 | 0 | -0.000353 | 0
1.1 | 11 | 11 | 2e-05 | -0.00048538 | 7.2e-05 | 8.4 | -0.00037065 | 17 | 8 | 0 | -0.000353 | 0
1.12 | 11.2 | 11.2002 | 0.00019 | -0.0004942 | -7.72e-05 | 8.6 | -0.00037948 | 17 | 8 | 0 | -0.000353 | 0
1.14 | 11.4 | 11.4003 | 0.00029 | -0.00050303 | -0.0001564 | 8.8 | -0.0003883 | 17 | 8 | 0 | -0.000353 | 0
1.16 | 11.6 | 11.6003 | 0.00026 | -0.00051185 | -0.000107 | 9 | -0.00039713 | 17 | 8 | 0 | -0.000353 | 0
1.18 | 11.8 | 11.8004 | 0.00036 | -0.00052068 | -0.000189 | 9.2 | -0.00040595 | 17 | 8 | 0 | -0.000353 | 0
1.2 | 12 | 12.0005 | 0.00046 | -0.0005295 | -0.000271 | 9.4 | -0.00041478 | 17 | 8 | 0 | -0.000353 | 0
[6 empty] | 9.6 | -0.0004236 | 17 | 8 | 0 | -0.000353 | 0
[6 empty] | 9.8 | -0.00043243 | 17 | 8 | 0 | -0.000353 | 0
[6 empty] | 10 | -0.00044125 | 17 | 8 | 0 | -0.000353 | 0
[6 empty] | 10.2 | -0.00045008 | 17 | 8 | 0 | -0.000353 | 0
[6 empty] | 10.4 | -0.0004589 | 17 | 8 | 0 | -0.000353 | 0
[6 empty] | 10.6 | -0.00046773 | 17 | 8 | 0 | -0.000353 | 0
[6 empty] | 10.8 | -0.00047655 | 17 | 8 | 0 | -0.000353 | 0
[6 empty] | 11 | -0.00048538 | 17 | 8 | 0 | -0.000353 | 0
[6 empty] | 11.2 | -0.0004942 | 17 | 8 | 0 | -0.000353 | 0
[6 empty] | 11.4 | -0.00050303 | 17 | 8 | 0 | -0.000353 | 0
[6 empty] | 11.6 | -0.00051185 | 17 | 8 | 0 | -0.000353 | 0
[6 empty] | 11.8 | -0.00052068 | 17 | 8 | 0 | -0.000353 | 0
[6 empty] | 12 | -0.0005295 | 17 | 8 | 0 | -0.000353 | 0
[6 empty] | 0 | #VALUE! | 0 | Nominal | 0 | Characterized Correction | 0
[6 empty] | 0 | #VALUE! | 0 | Nominal | 0 | Characterized Correction | 0

## Sheet2
magnification | 10 | x | (10 times the nominal size of the object)
Nominal Object size | 0.2 | inches
Magnified Object (N). | 2 | inches
Reading of magnified object | 2.01 | inches
magnification error, in inches | 0.01 | inches
magnification error of the actual object, in inches | 0.001 | inches
0.005
| 0.05 | error, expressed in power of x
| 10.05 | actual magnification
2.01
0.005
| 0.5 | % error of magnification
0.5